How to upgrade the firmware in XSO-S in the future:

Receive the latest firmware version from Bokar International

1.Copy file firmware XXX.zip and unpack this to one folder (All files must be in the samefolder)

2.Connect XSO-S (Data Store Connector) with PC (USB) via an X485-USB converter.NOTE: The X485-USB Converter needs an installed driver for proper operation.

3.When the converter is first connected to the PC, you should install the X485-USB driver.

4.Press and hold the enter button on the Oven

5.Turn the Power switch to ON position.

6.Release the “Enter” button. The oven display should show “Bootloader mode”

7.Run the program Flasher.exe on the PC (double click on the flasher icon)

8.The Flasher should find the serial port where the oven is connected.

9.The Flasher will ask “Do you want to start up update device?”,

10.Select “Yes” to confirm and press “enter”.

11.After the upgrade is finished, you should get the message “Write successful. Would you like torun application?”

12.Press “ENTER”

13.The application will start. The display in the oven should show X-KAR logo, name,serialnumber and firmware version number.

14.Flasher window will automatically close.